## Cover sheet
System / Subsystem | Openwater blood flow measurement device
Product(s): | Stroke detection Gen1A
Risk Analysis Type: | Design FMEA
Scope | measurement device
Intended Use reference |  | <- this specifies how the device is to be used (essential performance), and on whom? This helps determine, for instance, if a procedure delay is life threatening or benign
Risk Hazard Index Table (number-rev) | Dxxxxx
Hazard and Harms List (number-rev) | Dxxxxx

## Change Log
Date of Meeting/Update | Rev | Participants (Name(s) & Role(s)) | Description of Changes
| A |  | Quality Review
Reviewed/Generated content with core team

## DFMEA
Risk Analysis | [10 empty] | Risk Evaluation and Controls | [10 empty] | Pre to Post Mitigation Risk Hazard Index Delta | Evidence of Verification of Effectiveness | Risk arising from risk control measures | Labeling, IFU, Training, Use Error implications | Comments
ID # |  | Element: Component/ Sub-assy / Function | Part / Assy # | Sequence of Events |  |  |  |  | Clinical Harm | Person Subject to Harm
|  |  |  | Potential Failure Mode | Operation mode | Potential Cause of Failure | Hazardous Situation |  |  |  | Unmitigated Risk Level |  |  |  | Risk Control Measures |  |  | Residual Risk Level
| Subsystem | [5 empty] | Potential Functional Effects of Failure | Potential Effects on System Operation |  |  | Severity of Effect (S1-S4) | Occurrence (O1-O5) | Risk Hazard Index (O x S) | Risk Hazard Index Level | Inherent Safety by Design | Protective Measures | Evidence of Implementation | Severity of Effect (S1-S4) | Occurrence (O1-O5) | Risk Hazard Index (O x S) | Risk Hazard Index Level
[Unique identifier XYZ-001] | [What section of the product is being analyzed] | [Input from BOM, Intended/unintended function, and connections between components] | [Specific part or parts being referenced] | [Brief description of the failure, note where applicable: full, partial, intermittent, degraded, or unintentional failure ] | [Additional information] | [What causes this failure mode] | [What individual functions are affected by the failure] | [What effect does the failure have on entire system operation] | [impact to user, patient, operator, etc e.g. electric shock, burn, prolonged procedure, etc] | [patient, operator, etc] | [from column K, see Severity - Occurrence Table] | [Probability that the hazard would result in the harm without the mitigation in place, see Severity - Occurrence Table] | [Severity x Occurence, scale listed in "Severity - Ocurrence Table"] | [Risk Hazard Index Level] | [What requirements will be implemented to mitigate the risk? This should either justify no testing in column Q, or show that testing is needed] | [in the actual medical device and/or manufacturing process] |  |  | [Probability that the hazard would result in the harm with the mitigation in place] | [Severity x Occurence, scale listed in "Severity - Ocurrence Table"] | [Risk Hazard Index Level] | [For a significant delta,] |  | [Y/N - and explain] | [Does this risk indicate use error that can be mitigated by user instructions and training?] | [additional comments, etc. If new risks are identified, update the project's Hazards and Harms list]
FMEA-0001 | tower | AC Cable |  | Full power failure |  | - Hospital loss of power - Plug disconnected from socket by passerby. - Plug disconnected from socket by tower pulling away. | Entire system shuts down, data collection stops. | System inoperable | Procedure delay | N/A | S1 (Negligible) | O3 (Occasional) | 3 | Low | Long wand cable enables tower to be closer to wall; tower wheels lockable; durable cable sheathing; strain relief AC cord; locking cord AC-in socket; SSD instead of HDD | isolation transformer | Design requirements doc (wand, tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | -1 | AC cable will not detach from tower if pulled | Y - Wand cable could become a trip hazard | document "keep close to wall" in user manual
FMEA-0002 | tower | Power switch |  | switch complete failure, open or closed |  | faulty part | Loss of power to system or inability to shut off normally | fails open: system inoperable; fails closed: system can only be shut off by unplugging | Procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | high-grade part; recessed to avoid accidental contact; alternative shut-off switches for each subsystem | isolation transformer (with main power switch: all subsystems plug into this) | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | can bump tower switch and not accidentally shut system off (or turn it on) | N/A | N/A
FMEA-0003 | tower | Isolation transformer |  | current overlimit shut-off; internal failure; partial (per-socket) failure |  | overtaxed; faulty part | complete shutdown; loss of laser, AOM amp, PC or other subsystem; overheating | System inoperable | Procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | Toroid is a trusted manufacturer of component | N/A | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | N/A | N/A | N/A | OTS component
FMEA-0004 | tower | 120V AC power strip |  | complete or per-socket failure |  | faulty part | PC, fan or other subsystem shutdown (depends on failure mode) | System inoperable | Procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | quality part; includes add'l filtering | strip is inside locking tower doors | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | N/A | N/A |  | OTS component; octopus power outage means camera triggers still come out, but not high-voltage AOM signal
FMEA-0005 | tower | 12V Brick |  | Full electrical failure (burns out) |  | faulty part | octopus has no power | System inoperable | Procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | Reliable component | N/A | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | N/A | N/A | N/A | assuming that if octopus is unpowered, nothing can be triggered (TEST ME)
FMEA-0006 | tower | Fan(s) |  | Degradation burnout, or unintentional misalignment |  | faulty part or physical harm | no cooling, Laser & PC could shut down, potential component damage | system inoperable | Procedure Delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | redundancy of two fans | N/A | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | fan-disable test | N/A | N/A | OTS component
FMEA-0007 | tower | Fan controller |  | Partial or full fan controller failure |  | faulty part; electrical overload; wrong set point (human error) | no cooling, Laser & PC could shut down, potential component damage | system inoperable | Procedure Delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | no mechanical moving parts, less chance of failure | N/A | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | wrong set-point test | N/A | document correct temperature set-point and procedure in user manual | OTS component
FMEA-0008 | tower | 24V Brick for laser safety board |  | Full electrical failure (burns out) |  | faulty part or electrical overload | no power to laser safety board | system inoperable | Procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | Reliable component | N/A | Design requirements doc (laser safety board) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | pull brick, check that laser safety board shuts off power to laser | N/A | document in trouble shooting section | ots components
FMEA-0010 | PC | PC |  | Complete operational failure |  | bad power supply, memory or other hardware; bad sw upgrade; hacked | PC won't boot, won't run scanner SW, etc | system inoperable | procedure delay | N/A | S1 (Negligible) | O3 (Occasional) | 3 | Low | N/A | configuration management | run logs | S1 (Negligible) | O3 (Occasional) | 3 | Low | 0 | system diagnostics | N/A | document in diagnostic start-up and troubleshooting sections | OTS component w/custom SW; will be broken down further under SW rubric
FMEA-0011 | PC | PCIe expansion card |  | Partial or full failure of channels |  | faulty part(s) or physical damage during assembly | per-camera or all-camera failure | no data incoming | procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | SW diagnostics | connect cables to PC prior to shipping | cables are seated correctly | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | diagnostic output | N/A | document in diagnostic start-up and troubleshooting sections | OTS component
FMEA-0012 | PC | USB Connections |  | Full or partial failure of USB connection |  | Unreliable connection | per-camera failure | loss of data from one or more cameras | procedure delay | N/A | S1 (Negligible) | O5 (Frequent) | 5 | Low | two-cable connection from camera to PC | SW detection | diagnostic output | S1 (Negligible) | O5 (Frequent) | 5 | Low | 0 | diagnostic output | faulty SW | document in trouble shooting section
FMEA-0013 | tower | Thorlabs seed laser driver |  | hardware/electronics- electrical connections full or partial failure |  | internal or external electronics failure, or the connections were assembled poorly | laser stops emitting light, TA could get damaged | System cannot take measurements | procedure delay | N/A | S1 (Negligible) | O3 (Occasional) | 3 | Low | No relative motion between the source module and the power supply (both are bolted down), sequencer circuit to avoid turningg TA on with no/less current in seed driver. seed optical power checked before each scan. | secured within the tower, sequencer ckt | Design requirements doc (tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | - | N | Assembly instructions can point to securing plugging in the equipment
FMEA-0014 | tower | Thorlabs TA driver |  | hardware/electronics- electrical connections full or partial failure |  | internal or external electronics failure, or the connections were assembled poorly | laser stops emitting light | System cannot take measurements | procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | No relative motion between the source module and the power supply (both are bolted down), software sets and checks laser driver values before each scan, and data quality check intimates the error message for bad data | tower protects equipment from external tampering | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | - | N | Assembly instructions can point to securing plugging in the equipment
FMEA-0014 | tower | Thorlabs seed driver |  | software or communication full or partial failure |  | software changes/faulty USB Cable | laser cannot be turned on | System cannot take measurements | procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | No relative motion between the source module and the power supply (both are bolted down) | tower protects equipment from external tampering | Design requirements doc (tower) | S1 (Negligible) | O2 (Remote) | 2 | Low | 0 | - | N | Assembly instructions can point to securing plugging in the equipment
FMEA-0015 | tower | Thorlabs TA controller |  | electronics: not looking for right pulse width |  | internal settings are changed, but within the LSB limits | laser does not pulse at the right pulse width | System takes erroneous measurements, outputs more light than expected | procedure delay, minor skin burn, compromised data | N/A | S2 (Minor) | O3 (Occasional) | 6 | Low | LSB catches unexpected pulses (CW/high frequency/long pulses) | tower protects equipment from external tampering, S/W checks for parameter settings before starting scans. | Design requirements doc (tower), Design requirements doc (LSB) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -5 | - | N | Manual can indicate not to change any driver settings, tower locked at all times except servicing
FMEA-0016 | tower | Thorlabs TA controller |  | electronics: not looking for right frequency |  | internal settings are changed, but within the LSB limits + Octopus pulses at higher freq | laser does not pulse at the right frequency or pulse width | System takes erroneous measurements, outputs more light than expected | procedure delay, minor skin burn, compromised data | operator, patient (inconvenienced) | S2 (Minor) | O2 (Remote) | 4 | Low | LSB catches unexpected pulses (CW/high frequency/long pulses) | tower protects equipment from external tampering, S/W checks for parameter settings before starting scans. | Design requirements doc (tower), Design requirements doc (LSB) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -3 | - | N | Manual can indicate not to change any driver settings, tower locked at all times except servicing
FMEA-0017 | tower | Thorlabs TA controller |  | electronics: not set at the right TA Current |  | internal settings are changed, but within the LSB limits | laser outputs more power | System takes erroneous measurements, outputs more light than expected | procedure delay, minor skin burn, compromised data | operator, patient (inconvenienced) | S2 (Minor) | O2 (Remote) | 4 | Low | LSB catches higher energy pulses pulses | tower protects equipment from external tampering, S/W checks for parameter settings before starting scans. | Design requirements doc (tower), Design requirements doc (LSB) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -3 | - | N | Manual can indicate not to change any driver settings, tower locked at all times except servicing
FMEA-0017 | Octopus V1.0.0 | Octopus FX3/USB |  | Unable to communicate over USB with octopus [addressing full failure mode, if intermittant, can simply "try again"] |  | PCB/electronic damange, short circuit?, static?, physical damage to USB port, improper connection at assembly | can't program FPGA, no data can be acquired | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | Octopus inside protective case. If Octopus fails, Octopus can be physically replaced |  | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay.
FMEA-0018 | Octopus | GPIO (FSIN) | GPIO channels 1-4 BOTTOM | No output to FSIN channels [1 per camera], note: this row addresses full failure mode. intermittant failure will be addressed in software (missing frames ignored if 1 or 2, data retaken if many) |  | PCB/electronic damage (short circuit, static), physically breaking off the port | Loss of trigger to camera (will not acquire images) | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | Octopus inside protective case | if channel cannot be fixed, another channel will need to be used [up to 4 extra channels currently available], would require a minor software update, and physically moving the BNC cable | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay. | N/A
FMEA-0019 | Octopus | GPIO (FV) | GPIO channels 1-4 TOP | No output to FV (frame valid) channels [1 per camera], note: this row addresses full failure mode. intermittant failure will be addressed in software (missing frames ignored if 1 or 2, data retaken if many) |  | PCB/electronic damage (short circuit, static) | Loss of 'frame valid' signal to camera (will not process/save images) | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | Octopus inside protective case |  | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay. | N/A | if channel cannot be fixed, another channel will need to be used [up to 4 extra channels currently available], would require a minor software update, and physically moving the BNC cable
FMEA-0020 | Octopus | Digital Input | Digital In channel 1 bottom | No input from Wand On-button circuit to octopus |  | PCB/electronic damage (short circuit, static) | No input trigger to system, no pulses will occur, no data will be saved, SW will "hang" | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | Octopus inside protective case |  | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay. | N/A | There are 7 additional unused input channels, so user could manually move input to a different channel if the problem is channel specific (note this would require a minor
FMEA-0022 | Octopus | Digital Output 5 | TA QCW Trigger In | [intermittant failure] Intermittant loss of output to TA QCW Trigger therefore no laser light exits system |  | PCBA Component Damage, Unintended electrical connection(short) | Intermittant TA triggering resulting in some frames being dark. This will effect the quality of the output data (lower the mean intensity, etc), resulting in more failed data quality checks | Measurement taken is low quality, data quality check failed | Procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | Octopus inside protective case | There should be data quality checks in software to assure the variance in intensity between "bright" frames is low in case of this type of failure | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay. | N/A | up to 3 extra channels available, would require minor software update and physically switching SMA cable
FMEA-0024 | Octopus | FPGA |  | Inability to communicate with FPGA, bugs in firmware |  | PCB/electronic damage (short circuit, static) | Octopus cannot be programmed | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | Octopus inside protective case. If Octopus fails, Octopus can be physically replaced |  | Design requirements doc (wand, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 | N/A | Y - Difficulty accessing 'reset' button on octopus. Negligible harm - procedure delay. | N/A
FMEA-0025 | Octopus V1.0.0 | rework (resoldering) | Clock: SER4332CT-ND | Plug in USB, and octopus does not show up in device manager |  | bad solder re-work | Cannot communicate or program octopus | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | Octopus v1.0.1 will not require this rework |  | octopus v1.0.1 BOM | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1 | N/A | N | N/A
FMEA-0026 | Octopus V1.0.0 | rework (resoldering) | Memory IC: 24AA1026-I | Serial number won't show up |  | bad solder re-work | Cannot communicate or program octopus | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | Octopus v1.0.1 will not require this rework |  | octopus v1.0.1 BOM | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1 | N/A | N | N/A
FMEA-0027 | Detector Module | Trigger cables |  | Intermittent or full electrical connection failure |  | Poor soldering | Intermittant or no data collected | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | Cables have dual-layer heat shrink tubing, mechanical design includes strain relief |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1 | N/A | N | N/A
FMEA-0028 | Detector Module | USB cables |  | Intermittent or full electrical connection failure |  | poor connection | Intermittant or no data collected | System cannot take measurements | procedure delay | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | Cables verified during installation |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1 | N/A | N | N/A
FMEA-0029 | Detector Module | Cypress FX3 driver |  | Windows update prevents recognition of unsigned drivers |  | windows updates | Camera not recognized as "Openwater gumstick". Computer must be restarted in "advanced startup mode + F7 to allow unsigned drivers" | System cannot take measurements | procedure delay. | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | N/A | disable Windows 10 secure boot for duration of trial | can reboot normally and cameras respond | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | Y - minimal (attacker would have to somehow load an insecure driver); security by obscurity | N/A
FMEA-0030 | Detector Module | Himax register settings |  | Incorrect register settings on camera |  | poor communication with himax | variety of effects, from inoperative camera to bad data | Measurements taken may be noisy or invalid | Compromised data | Researcher | S1 (Negligible) | O3 (Occasional) | 3 | Low | Everything pre-tested in research system to avoid surprises |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -2 | N/A | N | N/A
FMEA-0031 | Source Module | Thorlabs seed laser |  | physical or electrical damage |  | human handling, electronics failure, degradation, fiber breakage | less or no light emitted, TA could be damaged | System cannot take measurements, or measurement data is compromised. | procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | sequencer ckt looks at current o/p of seed laser to turn TA off if below threshold, S/W queries seed optical power before each scan and turns TA on only after this check | Damping and protection provided by tower, fiber arranged carefully in racer track to avoid damage, placed out of easy reach. |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1 |  |  | Maintenance schedule | Have backup seed laser. In the future, implement read out from the butterfly mount tpo read OPTICAL power out of the seed laser to go directly to sequencer for continuous monotoring
FMEA-0032 | Source Module | TA |  | TA damaged |  | operator handling, electronics failure, physical damage | less or no light emitted | System cannot take measurements, or measurement data is compromised. | procedure delay | N/A | S1 (Negligible) | O2 (Remote) | 2 | Low | sequencer ckt looks at current o/p of seed laser to turn TA off if below threshold, S/W queries seed optical power before each scan and turns TA on only after this check | Damping the tower for mechanical shock, |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0033 | Source Module | Fiber Coupler |  | Bad/angled coupling to fiber |  | relative movement between TA and fiber coupler | beam coming out could be non uniform or not well coupled | Hot spots in the beam could lead to minor burns | skin damage | patient, operator | S2 (Minor) | O2 (Remote) | 4 | Low | source module and fiber coupler bolted down, and placed inside tower to avoid any mishandling and motion. diffusing spacers provided to expand the beam to reduce fluence and homogenize beam | Damping the tower for mechanical shock | Design requirements doc (laser safety board). | S2 (Minor) | O1 (Improbable) | 2 | Low | -2 |  | spacers not installed properly could fall? | Service and User manual to instruct service personnel to check beam shape with a beam profiler after every rework/realignment with the fiber.
FMEA-0034 | Source Module | Photodiode system (including pick off) |  | PD displacement, sensor damage, electronics, misalignment of splitter |  | operator handling, electronics failure | Photodiode sensing less or no light | Pulse delivered to patient is higher power than the system is aware of | minor skin burn | patient | S2 (Minor) | O2 (Remote) | 4 | Low | redundant mechanisms on laser safety board (electrical and optical monitring) | Laser safety goggles | Design requirements doc (laser safety board) | S2 (Minor) | O1 (Improbable) | 2 | Low | -2
FMEA-0035 | Source Module | Safety relay E-stop |  | switch fails |  | switch reliabiliy, physical trauma | TA does not turn off when the emergency button is depressed | System continues to emit light in an emergency | eye damage or skin damage | patient, operator | S2 (Minor) | O1 (Improbable) | 2 | Low | selecting relaible component | Laser safety goggles | Design requirements doc (laser safety board, tower) | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0036 | Source Module | Laser Safety board |  | LSB power supply failure |  | device failure | no power to laser | TA non-operational | procedure delay | patient, operator inconvenienced | S1 (Negligible) | O1 (Improbable) | 1 | Low | N/A (the laser safety board is a risk mitigation) | sw warns of loss of signal; this is potentially hard to troubleshoot as there is no indication of LSB status (service personnel will have to troubleshoot) | N/A | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 | N/A | N/A | should be included in Troubleshooting section of service manual
FMEA-0037 | Source Module | Laser Safety board |  | connection failure |  | mechanical impact | TA non-operational | TA non-operational | procedure delay | patient, operator inconvenienced | S1 (Negligible) | O1 (Improbable) | 1 | Low | N/A | sw warns of loss of signal; this is potentially hard to troubleshoot as there is no indication of LSB status (service personnel will have to troubleshoot) |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 | N/A | N/A | should be included in Troubleshooting section of service manual
FMEA-0038 | Source Module | Photodetector |  | dirty beam pick off |  | particulates | power to both photodetectors not correct | if the laser power output was higher than intended - we may not know | class 3R is the max power possible, possible eye damage | patient, operator | S2 (Minor) | O2 (Remote) | 4 | Low | good seal of source module, the optic is in another holder- minimal chance of accidental touching | Calibration measurement prior to each patient - would this catch a unusually high power? We should we add a step to measure the energy prior to each patient. LSB, SW setting would catch higher than normal TA drive current, so hard to be above known limits. |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -3
FMEA-0039 | Source Module | Photodetector/beam splitter |  | misaligned |  | tower impact | power to detector not correct | Pulse energy delivered to patient could be higher than indicated | class 3R is the max power possible, possible eye damage | patient, operator | S2 (Minor) | O2 (Remote) | 4 | Low | good seal of source module, the optic is in another holder- minimal chance of accidental touching | Calibration measurement prior to each patient - would this catch a unusually high power? We should we add a step to measure the energy prior to each patient. LSB, S/W settings would catch higher than normal TA drive current, so hard to be above known limits. |  | S2 (Minor) | O1 (Improbable) | 2 | Low | -2
FMEA-0040 | Source Module | Photodetector |  | damaged |  | power, impact or heat | power to detector not correct | Pulse energy delivered to patient could be higher than indicated | class 3R is the max power possible, possible eye damage | patient, operator | S2 (Minor) | O1 (Improbable) | 2 | Low | good seal of source module and keep clean in assembly |  |  | S2 (Minor) | O1 (Improbable) | 2 | Low | 0
FMEA-0041 | Source Module | TIA |  | damaged |  | electrical surge | Photodetector signal not correct | Pulse energy delivered to patient could be higher than indicated | class 3R is the max power possible, possible eye damage | patient, operator | S2 (Minor) | O1 (Improbable) | 2 | Low | electrical isolator used |  |  | S2 (Minor) | O1 (Improbable) | 2 | Low | 0
FMEA-0042 | Source Module | Photodetector |  | wire connection to safety board fails |  | wire breakage | TA stops getting triggers | No measurement can be taken | procedure delay. | patient | S1 (Negligible) | O2 (Remote) | 2 | Low | careful assemlby |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0043 | Source Module | Photodetector |  | Diode responsivity changes |  | temperature increase in device | Diode current reading to laser safety board is inaccurate | Pulse energy delivered to patient could be higher or lower than indicated | class 3R is the max power possible, possible eye damage | patient, operator | S2 (Minor) | O1 (Improbable) | 2 | Low | Can we calibrate the photodiodes at the temperature that we expect them to operate at when the device is being used? |  |  |  |  | #N/A | #N/A | #N/A
FMEA-0044 | Source Module | fiber connection between seed and TA |  | fiber gets damaged |  | general movement/jolt to the table, human carelessness (if it is exposed to outside) | seed power is not delivered to TA | TA gets damaged, no measurement can be taken | procedure delay. | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | SM is not easily accessible, fibers are placed on racetrack, almost no portion of fiber is exposed |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0045 | Source Module | Optical path - fiber coupler |  | fiber coupler gets misaligned/ comes off |  | general movement/jolt to the table, human carelessness (if it is exposed to outside) | no beam coming out of the wand so it gets into diagnostic mode, canot take measurements till this is solved | no measurements can be taken, potentially tougher troubleshoot | procedure delay. | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | cable covering the fiber and coupler |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0046 | Source Module | Optical path - fiber exit |  | Light back reflected into fiber |  | Wand held against reflective surface | Laser failure | System failure | procedure delay. | Patient | S1 (Negligible) | O1 (Improbable) | 1 | Low | isolator used, angled source fiber to prevent back reflection |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0047 | Wand | hand held |  | mechanical rigidity, short |  | dropped | hard to use | Handpiece could be difficult to use, measurements still possitble | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | add soft bumper, keep mass low |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0048 | Wand | Diffusing spacer tip |  | disposable fiber spacer to expand the beam falls off |  | dropped | MPE Exceeded | minor skin burn, data might be compromised | minor skin burn | patient | S2 (Minor) | O3 (Occasional) | 6 | Low | make the disposable tips well fitted, tegaderm used on top of the diffuser, makes it harder for the tip to fall off | protective eye wear to protect against laser radiation |  | S2 (Minor) | O2 (Remote) | 4 | Low | -2 |  |  | educate users during training and manual about putting these on, and checking for contact before measurements. | repeated inhouse measurements without the spacer tips have shown no damage with tissue
FMEA-0049 | Wand | Diffusing spacer tip |  | operator forgets to put diffusing spacer tip on |  | user error | MPE Exceeded | minor skin burn, data might be compromised | minor skin burn | patient | S2 (Minor) | O3 (Occasional) | 6 | Low | keep diffusing tips near wand, repeated reminders in the USER MANUAL, also in the scan UI before each scan | protective eye wear to protect against laser radiation |  | S2 (Minor) | O2 (Remote) | 4 | Low | -2 |  |  | educate users during training and manual about putting these on, and checking for contact before measurements. | repeated inhouse measurements without the spacer tips have shown no damage with tissue
FMEA-0050 | Wand | Diffusing spacer tip |  | variability in diffusing tip construction |  | not made properly | data consistency | erroneous measurement | erroneous measurement | operator, patient (inconvenienced) | S2 (Minor) | O1 (Improbable) | 2 | Low | have well defined construction and assembly instructions | protective eye wear to protect against laser radiation |  | S2 (Minor) | O1 (Improbable) | 2 | Low | 0 |  |  | educate users during training and manual about putting these on, and checking for contact before measurements.
FMEA-0051 | Wand | Fibers |  | fiber breaks |  | cable run over | signal from laser or to camera too low | no measurements can be taken | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | make robust sheathing |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0052 | Wand | Fibers |  | damaged fiber |  | cable run over | signal from laser or to camera too low | no measurements can be taken | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | make robust sheathing |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0053 | Wand | Fibers |  | dirty fiber face |  | not cleaned properly | signal from laser or to camera too low | erroneous measurement | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | cleaning procedure specified in the operator manual |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 |  |  | put cleaning in manual and lables
FMEA-0054 | Wand | springs |  | springs not attached correctly |  | dropped | no contact ot patient | erroneous measurement | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | robust attach of springs, with glue |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0055 | Wand | springs |  | springs stick |  | dirt within spring bearing | fiber is in retracted state,no contact ot patient | erroneous measurement | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O3 (Occasional) | 3 | Low | robust attach of springs, with glue |  |  | S1 (Negligible) | O2 (Remote) | 2 | Low | -1
FMEA-0056 | Wand | Wand trigger |  | short |  | switch mechanics failure | short mimics trigger on and octopus gate to TA is on | light is unintentionally emitted if a scan is running | eye damage | patient | S2 (Minor) | O3 (Occasional) | 6 | Low | Sturdier PCB, Roomier wand design to prevent fibers from rubbing up against the switch and triggering unintentionally |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -5 |  |  | instructions to wear laser safety goggles
FMEA-0057 | Wand | Wand trigger |  | short |  | switch mechanics failure | short mimics trigger off | no measurements can be taken | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O2 (Remote) | 2 | Low | Sturdier PCB, Roomier wand design to prevent fibers from rubbing up against the switch and triggering unintentionally |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | -1
FMEA-0058 | Wand | LED |  | LED failure |  | short | no feedback to operator | operator unsure what to do | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | select hi rel component, software provides redundant indication of data collection |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0059 | Wand | fibers |  | disconnected fiber |  | pulled too hard on wand/cable | no light emitted from disconnected fiber | erroneous measurement | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | extra strain relief, SW indicates bad measurement |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0060 | Wand | fibers |  | fibers exposed |  | cable damage | signal from laser or to camera too low | no measurements can be taken | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | robust sheath |  |  | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0
FMEA-0061 | tower, UI | keyboard, trackpad, mouse |  | incommunicative / blank |  | physical damage (trauma, spills, etc) | system inoperable | no measurements can be taken | delay in procedure | operator, patient (inconvenienced) | S1 (Negligible) | O1 (Improbable) | 1 | Low | components specifically intended for medical environments |  | physical inspection | S1 (Negligible) | O1 (Improbable) | 1 | Low | 0 | spill test | N/A | instructions for cleaning
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A
[13 empty] | #N/A | [8 empty] | #N/A

## Intended UseMisuse
| Sequence of Events |  |  |  | Clinical Harm | Person Subject to Harm | Risk Evaluation and Controls
Hazard | Forseeable Sequence of Events | Operation mode | Intended use, Use error, or Abnormal use? | Hazardous Situation |  |  | Unmitigated Risk Level |  |  |  | Risk Control Measures |  |  | Residual Risk Level
[7 empty] | Severity of Effect (S1-S4) | Occurrence (O1-O5) |  | Risk Level (O & S) | Inherent Safety by Design | Protective Measures | Evidence of Implementation | Severity of Effect (S1-S4) | Occurrence (O1-O5) |  | Risk Level (O & S) | Evidence of Verification of Effectiveness | Risk arising from risk control measures | Labeling, IFU, Training, Use Error implications | Comments
| [List the steps that could lead to the Hazardous Situation] | [Additional information] | [List all that apply] | [circumstance where people, property, and/or the environment is exposed to one or more potential source of harm (hazard)] | [impact to user, patient, operator, environment etc e.g. electric shock, burn, prolonged procedure, etc] | [user, patient, operator, etc] | [from column F] | [Probability that the hazard would result in the harm without the mitigation in place] |  | [Risk Hazard Index] | [What requirements will be implemented to mitigate the risk?] | [in the actual medical device and/or manufacturing process] |  |  | [Probability that the hazard would result in the harm with the mitigation in place] |  | [Risk Hazard Index] |  | [Y/N - and explain] | [Does this risk indicate use error that can be mitigated by user instructions and training?] | [additional comments, etc. If new risks are identified, update the project's Hazards and Harms list]
Mechanical Energy (gravity) | 1) Operator loses grip on the wand 2) Wand accelerates towards the ground |  | Use error | Wand hits patient (face, shoulder or lap) or operator (foot) on its way to the ground | Contusion, abrasion | Operator or patient | S2 (Minor) | O3 (Occasional) | 6 | Low | Elimination of sharp edges on wand design, weight of wand + height of normal use = low force on impact, ergonomic design of handle, non slip material |  | wand design | S2 (Minor) | O3 (Occasional) | 6 | Low |  | N | statement in manual about firmly gripping wand | Add: (Contusion, Abrasion) to harms list
Mechanical Energy (pinch point?) | 1) Operator hastily closes tower door |  | Use error | Operator skin of hand is pinched or finger is caught between door and frame | Contusion | Operator | S2 (Minor) | O2 (Remote) | 4 | Low | tower door closes manually, avoiding accidental pinching. Seal on tower door is relaxed with rounded edges, decreasing severity of potential harm. tower doors are not intended for day to day use, only for some servicing operations. |  | tower design | S2 (Minor) | O2 (Remote) | 4 | Low
Mechanical Energy (intertial force) | 1) Operator moves the tower around corners too fast |  | Use error | Operator gets hit by the monitor swinging due to inertia | Contusion | Operator | S2 (Minor) | O2 (Remote) | 4 | Low | Monitor's range of motion is restricted to allow less motion/swing. |  | tower design | S2 (Minor) | O2 (Remote) | 4 | Low |  |  | Operator instructed to move tower carefully
Bad Data | 1) Operator doesn't have wand pressed against patient for full duration 2) Operator does not hold wand steadily for full duration 3) One or more fibers do not make good contact with scalp |  | Use error | Poor data that will result in inaccurate result/diagnosis | Inaccurate diagnosis | Patient | S1 (Negligible) | O3 (Occasional) | 3 | Low | Software rejects the obviously poor data (like where there is no light, or there are missing frames) and requires user to retake. | Clear user instructions. | software | S1 (Negligible) | O1 (Improbable) | 1 | Low |  | None | Clear instructions | For initial studies we will inspect all raw data (csv files) manually.
Bad/Incomplete Data | 1) Operator takes hand off trigger before acquisition finishes too many times in a row |  | Use error | Poor data that will result in incomplete data set | Inaccurate diagnosis | Patient | S1 (Negligible) | O2 (Remote) | 2 | Low | Software rejects poor data and requires user to retake. | Clear user instructions. | software | S1 (Negligible) | O1 (Improbable) | 1 | Low |  | None | Clear instructions | For initial studies we will inspect all raw data (csv files) manually.
Mechanical energy (tripping hazard) | 1) Operator does not place tower near wall 2) Cord from tower to wall is an exposed tripping hazard |  | Use error | User or bystander trips and falls |  |  | S2 (Minor) | O2 (Remote) | 4 | Low |  |  |  | S2 (Minor) | O2 (Remote) | 4 | Low |  |  | Operator manual instructions to place tower near the wall
Chemical (biological contamination) | 1) Operator forgets to disinfect the wand or goggles |  | Use error | Spread of germs from user to user, or patient to use, etc |  | user/operator and patient | S3 (Serious) | O3 (Occasional) | 9 | Medium |  |  |  | S3 (Serious) | O3 (Occasional) | 9 | Medium |  |  | Instructions for disinfecting between uses, using tegaderm for additional insulation
Functional (erroneous data) | Operator releases trigger before enough frames are acquired during calibration |  | Use error | Data not taken | erroneous measurement |  | S1 (Negligible) | O4 (Probable) | 4 | Low |  | SW requires calibration measurement, which should diagnose this issue |  |  |  | #N/A | #N/A
Calibration error | Wand not in puck |  | Abnormal use | Bad Data | erroneous measurement |  | S1 (Negligible) | O3 (Occasional) | 3 | Low |  | SW requires calibration measurement, which should diagnose this issue |  | S1 (Negligible) | O1 (Improbable) | 1 | Low |  |  | user manual instructs on how tro take calibration data
Off-label use of PC | Someone uses PC for non-scan purposes, inadvertently disabling scanner in some way |  | Abnormal use | Scanner will not run, or runs incorrectly | Procedure delay | Operator or patient | S1 (Negligible) | O2 (remote) | 2 | Low | System checks SW config. Log to file and refuse to run if wrong. | Check log files. | sw test; logs | S1 (Negligible) | O2 (Remote) | 2 | Low
Data not backed up | 1) User forgets to backup data 2) PC disk crash |  | Use error | Data is lost | time wastage | Operator or patient | S1 (Negligible) | O3 (occasional) | 3 | Low | Reminder at the end of data collection/Automate data backup? |  |  | S1 (Negligible) | O3 (Occasional) | 3 | Low |  |  | Operator manual instructs operator to backup the data often
Hair between fiber and scalp | 1) Operator does not ensure good contact between scalp and fibers 2) Operator continues to collect data |  | use error | Erroneous data | delaying measurement | none | S1 (Negligible) | O4 (Probable) | 4 | Low | No metric set for evaluating the validity of the data we collected |  | SW design | S1 (Negligible) | O3 (Occasional) | 3 | Low | collecting reasonable signal | longer measurements time | Operator manual instructs operator to make good contact
Functional (erroneous Data) | 1) Operator places wand on incorrect placement, as described by software 2) Operator takes a measurements |  | Use error | Erroneous data | wrong interpretation and delayed measurement | patient and operator | S1 (Negligible) | O4 (probable) | 4 | Low | Clear graphics and instructions from SW to instruct operator which location to scan |  | SW design | S1 (Negligible) | O3 (Occasional) | 3 | Low | collecting reasonable signal | longer measurements time | Training procedure
Electromagnetic energy (near infrared light) | 1- operator starts a scan and depresses wand trigger 2 - operator gazes into laser beam |  | Abnormal use | Laser shines into operator's eye | eye damage | operator | S3 (Serious) | O1 (improbable) | 3 | Low | use of diffuser on the fiber | wears glasses |  | S2 (Minor) | O1 (Improbable) | 2 | Low |  |  | Laser label warning, safety training, operator manual warnings
Functional (erroneous Data) | 1) operator takes measurement 2) operator releases wand trigger before measurement completes |  | use error | no data taken | delay | NA | S1 (Negligible) | O4 (probable) | 4 | Low | Wand and software provides visual feedback that the operator can release the trigger |  |  | S1 (Negligible) | O3 (Occasional) | 3 | Low |  |  | training
Electromagnetic energy (near infrared light) | 1- Operator forgets to put on the fiber spacer tip. 2 -operator starts a scan and depresses wand trigger |  | use error | minor skin burn | minor skin burn | patient | S2 (Minor) | O2 (Remote) | 4 | Low | Operator manual instructs user to ensure the tip is fitted before starting the scan | wears glasses |  | S2 (Minor) | O1 (Improbable) | 2 | Low |  |  | Laser label warning, safety training, operator manual warnings
Bad data | 1- operator does not insert the diffusing spacer properly 2 -diffusing spacer is unseated 3 - Operator continues taking the scan |  | use error | bad or inconsistent data | delay | operator | S2 (Minor) | O1 (improbable) | 2 | Low | Operator manual instructs user to ensure the tip is fitted before starting the scan | wears glasses |  | S2 (Minor) | O1 (Improbable) | 2 | Low |  |  | operator manual, and training will instruct on how to use the diffusing spacer

## Severity - Occurrence Table
Risk Hazard Index Table |  |  | Severity
|  |  | S1 | S2 | S3 | S4 | Risk Hazard Index
|  |  | Negligible | Minor | Serious | Catastrophic | High (14-20) | Unacceptable Risk, requires mitigation
Occurrence | O5 | Frequent | 5 - Low | 10 - Medium | 15 - High | 20 - High | Medium (7-13) | Requires mitigation or justification
| O4 | Probable | 4 - Low | 8 - Medium | 12 - Medium | 16 - High | Low (1-6) | Acceptable Risk
| O3 | Occasional | 3 - Low | 6- Low | 9 - Medium | 12 - Medium
| O2 | Remote | 2 - Low | 4 - Low | 6 - Low | 8 - Medium
| O1 | Improbable | 1 - Low | 2 - Low | 3- Low | 4 - Low
Definitions
Occurrence
O5 | Frequent | >10%
O4 | Probable | <= 10%
O3 | Occasional | <= 1%
O2 | Remote | <= 0.1%
O1 | Improbable | <=0.01%
Severity
S4 | Catastrophic | An effect that likely results in death, or major environmental harm
S3 | Serious | Injury or illness that is one or more of the following: (1) Is life-threatening (2) Results in permanent impairment of a body function or permanent damage to a body structure (3) Necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure (4) Causes Environmental Harm requiring specialized remediation
S2 | Minor | An effect without significant consequence which does NOT not necessitate medical, or surgical to preclude permanent impairment of a body function or permanent damage to a body structure.
S1 | Negligible | An effect that does not require therapy or result in any consequences.
Risk Hazard Index
High | Unacceptable Risk, requires mitigation
Medium | Requires mitigation or justification
Low | Acceptable Risk

## Clinical Harm Table
Harms | Definition | Severity | Severity Rating
Prodecure Delay | Measurement procedure takes longer than intended | Negligible | S1
Compromised data?
Minor skin burn
Contusion |  | Minor | S2
Abrasion | superficial damage to the epidermis | Minor | S2
Serious skin burn | permanent damage to the epidermis | Serious | S3
Eye damage | permanent damage to eye | Serious | S3

## DFMEA Process
DFMEA stands for Design Failure Mode and Effect Analysis
It serves to review designs to identify what might go wrong, in which way would the failure manifest, and what would be the consequences of those failures.
STEP 1: REVIEW THE DESIGN | Columns A-D
Multi-disciplinary team identifies every system, subsystem, interface, and component.
STEP 2: IDENTIFY FAILURE MODES | Columns E-F
Once systems/components are decoupled, the next step is to thoroughly assess each to determine all the possible ways they can fail.
STEP 3: DEFINE THE CAUSE/MECHANISMS OF FAILURE | Column G
Conciseness and completeness are the keywords at this stage. You don’t want to get lost in long descriptions where it takes a lot of time to understand what the causes or mechanisms are. Every cause or mechanism you can think of should be listed, but staying focused with pertinent ones will ensure you don’t go off track.
STEP 4: LIST THE POTENTIAL EFFECTS OF EACH FAILURE MODE | Columns H-I
Functional effects - individual functions that are affected
Effects on system operation - what would be the effect on the entire system? How would the operation or output of the system change?
STEP 5: IDENTIFY THE RESULTING HARM | Column J-K
What would be the resulting impact, and to whom would it occur? - this should be evaluated in the context of customer use
Could be bodily harm, environmental harm, delayed procedure, etc..
Add new Harms in Column J to Clinical Harm Table
STEP 6: ASSIGN RISK HAZARD INDEX | Columns L-O
Severity of effect & Probability of occurrence = Risk Hazard Index (Severity-Occurence Table)
STEP 7: ANALYSIS OF RISK CONTROL OPTIONS | Columns O-Q
Inherent safety by design: how will the device be designed to avoid the hazardous situation? - this should always be the first choice to deal with a risk
What design decisions mitigate the likelihood of occurrence? - second choice
Evidence of implementation - should point to a design or requirements document that includes the above design controls
STEP 8: ASSIGN (POST MITIGATED) RISK HAZARD INDEX | Columns R-T
STEP 9: EVIDENCE OF MITIGATION EFFECTIVENESS | Column U
Usually done by Quality Assurance / V&V
STEP 10: IDENTIFY ANY ADDITIONAL RISKS ARISING FROM CONTROL MEASURES
Did our risk control measures cause a resulting additional risk
STEP 11: IDENTIFY ANY ADDITIONAL RISK CONTROL MEASURES (SAFETY INFORMATION

## Intended UseMisuse Process
Walk through all the steps required for use of the product. For each step:
STEP 1: IDENTIFY POTENTIAL HARMS (HAZARDS) that could arise from: | COLUMN A
1) Intended use - “use for which a product, process or service is intended according to the specifications, instructions and information provided by the manufacturer”
2) Use error - unintentional mistakes that can happen during use
3) Abnormal use - intentional acts of misuse, such as not following directions, and/or using the medical device for a different off-label purpose
Some examples of hazards (potential harms):
Electromagnetic Energy
- Line Voltage
- Leakage Current
- Electric Fields
- Ionizing Radiation
Thermal Energy
Mechanical Energy
- Gravity
- Vibration
Biological
- Bacteria
- Viruses
Chemical
Function
Operational
REMEMBER: Hazards (potential harms) can affect people, property, OR environment
Steps to use the product can be found in the Operator Manual
STEP 2: LIST THE SEQUENCE OF EVENTS | COLUMNS B-D
What sequence of events would lead to the circumstances where the hazard could take place?
Note if these events would occur in a specific operating mode (column C) and whether the events fall under: intended use, use error, or abnormal use (column D)
STEP 3: WHAT IS THE HAZARDOUS SITUATION | COLUMN E
The circumstances where the hazard could take place
STEP 4: WHAT IS THE HARM THAT COULD OCCUR? | COLUMN F
STEP 5: TO WHOM COULD THE HARM OCCUR? | COLUMN G
STEP 6: ASSIGN RISK HAZARD INDEX | Columns H-J
Severity of effect & Probability of occurrence = Risk Hazard Index (Severity-Occurence Table)
STEP 7: ANALYSIS OF RISK CONTROL OPTIONS | Columns K-M
Inherent safety by design: how will the device be designed to avoid the hazardous situation? - first choice
What design decisions mitigate the likelihood of occurrence? - second choice
Evidence of implementation - should point to a design or requirements document that includes the above design controls
STEP 8: ASSIGN (POST MITIGATED) RISK HAZARD INDEX | Columns N-P
STEP 9: EVIDENCE OF MITIGATION EFFECTIVENESS | Column Q
Usually done by Quality Assurance / V&V
STEP 10: IDENTIFY ANY ADDITIONAL RISKS ARISING FROM CONTROL MEASURES
Did our risk control measures cause a resulting additional risk
STEP 11: IDENTIFY ANY ADDITIONAL RISK CONTROL MEASURES (SAFETY INFORMATION)

## Notes
Risks/Failures that have already come up